## Sheet1
	Measurements taken from TLU front panel (mV) for indicated set value (mV)					Errors
PMT Chan	100	200	300	400	500	600	700	800	900	Gain	Offset
0	100.3	200.2	300.4	401.0	500.5	600.8	701.6	801.0	901.0	1.0012	0.1389
1	92.1	192.0	291.3	391.8	490.6	589.8	689.8	789.1	889.3	0.9959	-7.2806
2	97.9	197.2	296.5	398.0	497.5	597.1	697.6	797.0	897.1	0.9996	-2.4694
3	104.2	203.4	303.6	404.1	504.0	603.2	703.5	803.1	902.5	0.9985	4.2444

Vref (mV)=	1000	# bits =	10	Vlsb (mV)=	0.9766

	Instructions: / Make sure the Vref value is correct (an unmodified TLU will be 1000mV, and / if LC1 is cut and LO1 is jumpered it will be 2000mV). The number of bits for a / standard TLU’s Vcntl DAC is 10 (it uses the AD5316 chip). The Vlsb (the / smallest difference between two DAC bit codes) is calculated. Once calibration / is done, the error should generally be no more than +/-1/2Vlsb. Once the / above information is filled-in/correct, then for each of the four PMT channels, take / the nine measurements per channel by specifying that voltage in the EUDAQ / configuration file, or by using the TLUControl.exe program. To use the EUDAQ / configuration file, make a copy of a working configuration file to use while filling / in this spreadsheet, ensure that no individual PMTs have a specified control / voltage set (i.e. make sure that PMTVcntl1 through PMTVcntl4 are not present / in the configuration file), set the PMTVcntl parameter to the value desired to fill / in a column in the above spreadsheet (i.e. PMTVcntl = 100 ... it will set all the / PMTs to the same control voltage). Measure each of the four PMT channels’ / PMT control voltages on the front panel LEMO connector (pin 3) for each of the / control voltage settings in the spreadsheet (this is setting the voltage 9 times / and taking 4 x 9 = 36 measurements, but it should only need to be done once / or at most when calibration is absolutely required for an experiment, so don’t / be overwhelmed at the prospect). Mark down the measured values in the / matching row and column of the spreadsheet. To use the TLUControl.exe program, / just use the “TLUControl.exe -q -pv <control voltage value>” invocation for each / of the nine columns (100 through 900 in steps of 100mV), and enter the measured / values into the spreadsheet. The Gain Error and Offset Error will be calculated / for each of the PMT control voltage channels. Enter these 8 values into all of your / configuration files (they can’t be used with the TLUControl.exe program) using the / PMTGainError1 through PMTGainError4, and PMTOffsetError1 through / PMTOffsetError4 parameters. These values will be applied to correct the DAC.